## Slide 1
Chart: Nonconformities Q3 n=90
| Lab | Nonlab
 | 47 | 53
Chart: Nonconformities YTD n=250
| Lab | Nonlab
 | 42 | 58
 | Nonlab |  | Lab | 
 | Q3 | YTD | Q3 | YTD
Pre-exam | 91% | 93% | 50% | 65%
Exam | 0% | 0% | 25% | 12%
Post-exam | 9% | 7% | 25% | 23%
Cape Clinic Hospital Laboratory
Quality Report – Q3
05/09/2016
IV. Nonconformities (continued)
QGen-MR-F004;v2
06/13/2016
Page 3 of 6
V. Internal Audit
Q3 – 3rd Quarter Interim YTD = Year to Date
Review summary of internal audits performed since last management review (Interim Q2) | 
Audit Date: 31/07/2016 | Process Audited: Document Control – horizontal audit
Good News | 
Master File Index is current. Master Files are current and well organized. | 
Nonconformances | 
Out of a sampling of 30 SOPs, 6 locations did not have the current version and 11 locations had an obsolete version available which is a nonconformity against ISO15189 Clause 4.3 regarding document control measures ensuring that only the current and approved version is present and obsolete versions are removed. Post-it notes were used to revise approved documents at the workbench which is a nonconformity against QGen-DocCon-P001 and ISO15189 Clause 4.3e regarding amendments by hand. | 
Opportunities for Improvement | 
Staff expressed frustration with management requiring the use of approved documents. 5 of 8 staff members voiced frustration with post-it notes falling out and onto the floor, inability to find the document, or which one to use. All 8 voiced that they do not have confidence in the document control system, especially now that they must sign attestation sheets. | 
Audit Date: 18/06/2016 | Process Audited: Following the Total Testing Process –vertical audit
Good News | 
Handout: Quality Report 4-81

## Slide 2
Cape Clinic Hospital Laboratory
Quality Report – Q3
05/09/2016
QGen-MR-F004;v2
06/13/2016
Page 5 of 6
Resource Status Report | Phlebotomy/ Reception | Chemistry | Hematology | Microbiology | Serology | Blood Bank | Molecular Biology | Histology | Cytology
Competency - Determine the % of individuals >30 days overdue: |  |  | 4/4 = 100% |  |  |  |  |  | 
Safety Training –Determine % any individuals >30 days overdue: |  |  | 4/4 = 100% |  |  |  |  |  | 
Calibrations – Identify any equipment calibrations >30 days overdue: | NA | 1/33 (GGT); 3% | FBC 15/15 = 100% |  |  |  | 1/1 VL kits; 100% | NA | NA
Scheduled Maintenance. Identify any equipment services >30 days overdue: | NA |  | FBC and CD4 analyzers; 3 micro-scopes | 3 micro-scopes |  | 2 micro-scopes |  | 4 micro-scopes | 2 micro-scopes
……….. |  |  |  |  |  |  |  |  | 
……….. |  |  |  |  |  |  |  |  |