Validation Summary Report

Purpose: (Validation (Re-Validation   X Other: Verification of Manufacturer’s Claims

Description of Equipment/Process:

Equipment/Process:  Roche Cobas c501 Chemistry Analyzer regarding AST only

Serial Number: 12345

Location: Cape Clinic Laboratory, 111 Main Street, Arusha, Tanzania

Date: November 1 through 30, 2013

FDA Approval Status:  X  Approved  (Not approved

Procedure:

Refer to the Cape Clinic Laboratory Validation Plan for Cobas c501 regarding AST

Results:  All raw data reports and statistical analysis can be found in the Cobas c501 Validation binder.

- Precision -- refer to tab A

| Analyte | Expected Results | Observed Results | Acceptability |
| --- | --- | --- | --- |
|  |  | Between Day |  |
|  | Manufacturer’s Precision | 33% of CLIA | Normal Control CV% | Abn Control CV% |  |
| Ex:ALT | 3.3% | 6.6% | 3.8% | 4.3% | Acceptable |
| AST | 1.3% | 6.6% |  |  |  |

| Analyte | Expected Results | Observed Results | Acceptability |
| --- | --- | --- | --- |
|  |  | Within Run |  |
|  | Manufacturer’s Precision | 25% of CLIA | Normal Control CV% | Abn Control CV% |  |
| Ex:ALT | 2.6% | 5% | 2.6% | 4.3% | Acceptable |
| AST | 0.8% | 5% |  |  |  |

- Accuracy -- refer to tab B

Correlation with Cape Clinic chemistry department’s XYZ Chemistry Analyzer using forty patient specimens spanning the AMR for AST will be analyzed in duplicate on each instrument. At least twenty of the specimens selected must be outside of the current reference interval for AST.  Specimens selected will be analyzed on the same day of collection and within two hours between each instrument.  Replicates will be analyzed in different runs and then reviewed for any significant discrepancies, such as sample mix-ups or transposition errors.

Patient samples will be collected and handled in accordance with Cape Clinic’s collection and handling procedures.  A serum aliquot from a centrifuged red top collection tube will be analyzed. Sufficient quantities of each sample (minimum 1.0 ml serum) must be available to: 1-) analyze duplicates by the test methods, 2-) analyze duplicates by the comparative methods, and 3-) perform follow-up studies if needed.

Data collection will occur over a span of 10 working days, analyzing four specimens per day.

| Analyte | Total Allowable Error | Correlation Coefficient (r) | Linear Regression Statistics | Acceptability For Linear Regression |
| --- | --- | --- | --- | --- |
|  |  | Expected >0.975 | Slope | Intercept |  |
| Ex.ALT | 0.5U/L or 20% | 0.997 | 1.002 | 2.19 | Acceptable |
| AST | 20% |  |  |  |  |

| Analyte | Xc | Linear Regression Equation | Bias at Xc | %Bias at Xc |
| --- | --- | --- | --- | --- |
| Ex.ALT | 40 U/L | Y= 1.002 X + 2.19 | 2.3 U/L | 5.7% |
| AST | 35 U/L |  |  |  |
| AST | 120 U/L |  |  |  |

- Linearity -- refer to tab C

  - Linearity verification will be determined using patient samples previously analyzed on the XYZ analyzer.  A known high near or slightly above the upper measurement range and a known low near the limit of detection will be used.

| Analyte | Linear Regression Statistics | Allowable Systematic Error | Linear Range Verified | Evaluation |
| --- | --- | --- | --- | --- |
|  | Slope | Intercept | 50% of TEa |  |  |
| Ex:ALT | 0.970 | 0.282 | 10% | 2.5-770 | Linear |

- Analytical Measurement Range (AMR) and Clinical Reportable Range (CRR) -- refer to tab C

| Analyte | Mfg’s AMR | Low Value Verified | High Value Verified | Reportable Range | Dilutions | CRR |
| --- | --- | --- | --- | --- | --- | --- |
| Ex:ALT | 5-700 U/L | 2.5 | 770 | 5-700 | 1:10 | 5-7000 |

- Sensitivity and Specificity -- refer to tab D for instrument package inserts

| Summary of Manufacturer’s Claims for Sensitivity and Specificity |
| --- |
| Analyte | Specificity (Interfering Substances) | Sensitivity |
| AST | Icterus – No significant effect Hemolysis – No significant effect Lipemia – No significant effect; >Abs flagging may occur Drugs – No interference at therapeutic concentrations; Isoniazid may cause artificially low and Furosemide may cause artificially high AST results at therapeutic concentrations | 5 U/L |

- Reference ranges -- refer to tab E

  - The reference range studies have been verified for the following populations:

    - Adult reference ranges –Reference ranges for males and females were transferred from the XYZ manufacturer’s expected values.  The transference was verified using forty serum samples (twenty adult males and twenty adult females) from Cape Clinic’s employee population.

    - Pediatric reference ranges– Reference ranges for pediatrics were adopted from Ugandan reference ranges in use at the XYZ Research lab in Kampala, Uganda.  The ranges were evaluated by the Cape Clinic medical leadership team in consultation with local pediatricians and determined to be appropriate for the local Tanzanian population. Adoption of these ranges was approved by the Cape Clinic Medical Director.  All ranges will be verified and monitored over time as appropriate data becomes available.

| Analyte | Adult Reference Ranges | Reference Range Cited | % Verified (Expected ≥90%) |
| --- | --- | --- | --- |
| AST | Males ≤ 38 U/L Females ≤32 U/L | XYZ Manufacturer Range |  |

Reference Range Approval

Medical Director: _______________________________  Date: _______________

Insert Medical director name here

Acceptability of Method

The Cobas c501 Chemistry Analyzer may be judged acceptable for the AST analyte if one of the following conditions is met

- Manufacturer’s claims for linearity, precision and accuracy have been verified

- The total error calculation (bias + 3 sd or %bias + 3 %CV) for the test method is less than the CLIA total allowable error for each Medical Decision Point (Xc).

- The Sigma-metric is > 3.0.

| Concentration of Xc | AST %TEA at concentration of Xc | Long-term Precision (%CV) | %Bias or %difference with comparative method | Total Error (in %) | Sigma Metric | Acceptability |
| --- | --- | --- | --- | --- | --- | --- |

Method Approval

Approved / Not Approved

If not approved, provide recommendations/corrective actions below.

Laboratory Director: ____________________________  Date: _________________

Insert Lab Director name here

Prepared by: ______________________________  Date:___________________

Insert name and title here