Chemistry Precision Guidelines

PRECISION is reproducibility -- the agreement of the measurements of replicate runs of the same sample.  Replication experiments are performed to estimate the imprecision or random error of a given analytical method.

- Short -Term (Within-run/Day)

  - Sample:

    - Two levels (Low / High or Normal / Abnormal)

    - Patient or quality control

    - Select values near the medical decision point(s) of interest for the analyte

  - Testing:

    - Ensure there is a sufficient reagent to perform all 20 tests.

    - Run each sample 20 times on the same run, if possible, or least within the same day.

  - Acceptability criteria:

    - Calculate the coefficient of variation (CV) for each level using 20 data points.

    - Compare the calculated CV to the manufacturer’s stated precision claims found in the package insert.

    - If manufacturer’s precision cannot be met, it is acceptable to attain precision that is <25% of the CLIA Allowable Error.

    - If Short -Term precision is unacceptable, consult the instrument’s manufacturer for technical assistance.

    - If unable to resolve issues with short-term precision, the method validation process should be discontinued and a new method selected for potential implementation.

- Long-Term (Between-run/Between Day)

  - Sample:

    - Two levels (Low/High or Normal/Abnormal)

    - Patient or control serum.  A lab may already have this data available from their daily QC runs.

    - Select values near the medical decision point(s) of interest for the analyte

  - Testing: Run each sample 1 time per day for 20 days for a minimum of 20 total data points for each level of material used.

  - Acceptability criteria:

    - Calculate the CV for each level using 20 data points

    - Compare to manufacturer’s stated precision claims found in the package insert.

    - If manufacturer’s precision cannot be met, it is acceptable to attain precision that is <33% of the CLIA Allowable Error

    - If Long-Term precision is unacceptable, consult the instrument’s manufacturer for technical assistance.

- References

  - GCLP Workshop and Workbook18-20 May 2008, page 16

  - Clinical and Laboratory Standards Institute (CLSI). User Verification of Performance for Precision and Trueness: Approved Guideline-Second Edition. CLSI document EP15-A2 (ISBN 1-56238-574-7). Clinical and Laboratory Standards Institutes, 940 West Valley Road, Suite 100, Wayne, Pennsylvania 19098-1898 USA, 2005.

  - NCCLS. (Currently CLSI) Evaluation of Precision Performance of Quantitative Measurement Methods; Approved Guideline—Second Edition. NCCLS document EP5-A2 (ISBN 1-56238-542-9). NCCLS, 940 West Valley Road, Suite 1400, Wayne, Pennsylvania 19087-1898 USA, 2004.

  - Westgard, James O. (2003). Basic Method Validation (2nd ed.). Madison: Westgard QC, Inc.

  - Westgard, James O. (2008). Basic Method Validation (3rd ed.). Madison: Westgard QC, Inc.